PySpectrometer2 Software Architecture and Design

Target Architecture and Detailed Design

Project: PySpectrometer2

Revision: 1.1

Revision Date: July 14, 2026

Status: Controlled Engineering Document

# Revision History

| Revision | Date | Description | Author |
| --- | --- | --- | --- |
| 1.0 | July 2026 | Initial document. | Project Team |
| 1.1 | July 14, 2026 | Added Raspberry Pi as a first-class supported platform and updated platform, hardware abstraction, camera interface, testing, deployment, and migration requirements. | Project Team |

# 1. Architectural Objectives

Create one scientific application core shared by Linux, Windows, and Raspberry Pi deployments.

Prevent any one operating system or hardware family from owning the architecture.

Preserve proven spectral algorithms while separating acquisition, processing, calibration, presentation, persistence, and diagnostics.

# 2. Architectural Principles

One application core; multiple platform and hardware adapters.

Dependency direction shall point inward toward platform-neutral scientific logic.

Camera acquisition shall be abstracted from spectrum extraction and analysis.

Configuration and logging shall be centralized.

Synthetic spectra shall use the normal analysis pipeline rather than a separate simplified code path.

# 3. Proposed Component Model

Application Core: startup, shutdown, orchestration, dependency construction, CLI dispatch.

Configuration Service: INI loading, validation, defaults, profile selection, runtime overrides.

Logging Service: centralized structured logs, verbosity levels, rotation.

Camera Abstraction Layer: common camera interface with USB, Linux, Windows, and Raspberry Pi camera adapters.

Acquisition Service: frame capture and acquisition metadata.

Spectrum Extraction: converts camera frames into one-dimensional spectra.

Calibration Service: wavelength calibration, persistence, validation, and application.

Scientific Analysis: peak detection, line fitting, FWHM measurement, SNR estimation, and related algorithms.

Synthetic Spectrum Generator: creates reproducible multi-line test spectra with configurable wavelength, amplitude, FWHM, SNR, baseline, sampling, and random seed.

Data Persistence: raw frames, spectra, metadata, calibration, and synthetic-generation parameters.

Presentation Layer: GUI and plotting without embedding scientific algorithms in interface code.

# 4. Platform Abstraction

Linux, Windows, and Raspberry Pi are equal first-class deployment targets.

Platform-specific device discovery, drivers, paths, services, and camera APIs shall remain behind adapters.

Raspberry Pi support may use native camera APIs or USB cameras without changing the scientific core.

Performance-sensitive behavior may be tuned per platform, but numerical meaning and data formats shall remain consistent.

# 5. Camera Interface Design

A camera adapter should expose initialization, capability discovery, configuration, frame acquisition, health/status, and clean shutdown.

Initial adapter families should include generic USB/OpenCV-compatible cameras and Raspberry Pi-native camera interfaces where supported.

Scientific modules shall consume normalized frames and metadata, not vendor-specific camera objects.

# 6. Synthetic Spectrum Generator Design

Inputs: one or more center wavelengths, amplitudes, FWHM values, wavelength bounds, sample spacing, SNR, baseline level, baseline slope, profile type, and random seed.

Initial profile: Gaussian. Extension points: Lorentzian and Voigt.

Output: the same spectrum data object accepted by the standard analysis pipeline.

Reproducibility: identical inputs and random seed shall produce identical generated data within platform-independent numerical tolerances.

Validation uses: peak recovery, FWHM recovery, SNR estimation, blended peaks, calibration regression, algorithm comparison, and edge-case testing.

# 7. Recommended Package Structure

src/pyspectrometer2/app/ — startup, CLI, orchestration

src/pyspectrometer2/config/ — INI schema, validation, profiles

src/pyspectrometer2/platform/ — platform-specific services

src/pyspectrometer2/cameras/ — camera interface and adapters

src/pyspectrometer2/acquisition/ — frame acquisition

src/pyspectrometer2/processing/ — spectrum extraction and preprocessing

src/pyspectrometer2/calibration/ — wavelength calibration

src/pyspectrometer2/analysis/ — scientific analysis

src/pyspectrometer2/synthetic/ — synthetic spectrum generation

src/pyspectrometer2/data/ — persistence and metadata

src/pyspectrometer2/ui/ — GUI and plotting

tests/ — unit, integration, regression, hardware, and platform tests

# 8. Testing Architecture

Unit tests for pure algorithms and configuration validation.

Golden-data regression tests preserving current proven behavior.

Synthetic-data tests with deterministic seeds.

Camera contract tests for every adapter.

Platform acceptance tests on Linux, Windows, and Raspberry Pi.

Hardware-in-the-loop tests where cameras or native interfaces are required.

# 9. Design Decision: Raspberry Pi

Raspberry Pi is restored and retained as a first-class platform because of its substantial DIY, educational, maker, and portable-instrumentation user base.

The architectural rule is support without coupling: the Pi receives full support, but the core application remains independent of Pi-specific APIs.