SpectrometerV2 Requirements and Features Document

Software Requirements Specification

Project: SpectrometerV2

Revision: 1.7

Revision Date: July 16, 2026

Status: Controlled Engineering Document

# Revision History

| Revision | Date | Description | Author |
| --- | --- | --- | --- |
| 1.0 | July 2026 | Initial document. | Project Team |
| 1.1 | July 14, 2026 | Added Raspberry Pi as a first-class supported platform and updated platform, hardware abstraction, camera interface, testing, deployment, and migration requirements. | Project Team |
| 1.2 | July 16, 2026 | Added implementation-status annotations to every requirement, reflecting SpectrometerV2 Sprint 1 (configuration and logging services complete; CLI, error handling, camera, and synthetic scaffolding in progress). Rebranded document project name to SpectrometerV2. | Project Team |
| 1.3 | July 16, 2026 | Synced status: TEST-002 (synthetic verification) now in progress via the Phase 1 golden and synthetic-recovery test suite. | Project Team |
| 1.4 | July 16, 2026 | Synced status: synthetic generator implemented - SYN-001..006, 008, 009 Implemented; SYN-007 In Progress (Gaussian done; Lorentzian/Voigt future). | Project Team |
| 1.5 | July 16, 2026 | Synced status: camera abstraction implemented - HW-001..004 and PLAT-005 Implemented (USB, Raspberry Pi, and synthetic adapters behind the common contract). | Project Team |
| 1.6 | July 16, 2026 | Synced status (Phase 6): CLI-004 and ERR-001 Implemented; PLAT-001/002/003/007 In Progress; PLAT-004 Implemented; DATA-001 In Progress; DATA-002 Implemented. | Project Team |
| 1.7 | July 16, 2026 | Synced status (Phase 8 groundwork): TEST-001 and ACC-001 In Progress - installable package and a platform acceptance smoke test validate the full stack on the current platform (Linux); Windows/Pi pending. | Project Team |

Implementation status (Rev 1.2): [Implemented] = built and verified in the SpectrometerV2 codebase; [In Progress] = partially implemented or scaffolding/interface in place; [Planned] = not yet started. Reflects branch sprint-1-skeleton as of July 16, 2026.

# 1. Purpose

Define the functional and non-functional requirements for the modernization of SpectrometerV2 into a maintainable, modular, cross-platform scientific instrument application.

# 2. Design Philosophy

Think like a physicist. Build like an engineer.

The system shall preserve accessibility for DIY builders while supporting professional engineering and laboratory environments.

# 3. Platform Requirements

PLAT-001: The system shall fully support Linux as a first-class platform.  [In Progress]

PLAT-002: The system shall fully support Windows as a first-class platform.  [In Progress]

PLAT-003: The system shall fully support Raspberry Pi as a first-class platform.  [In Progress]

PLAT-004: The core application architecture shall not require Raspberry Pi-specific hardware or software. Raspberry Pi-specific functionality shall be isolated behind clearly defined platform and hardware interfaces.  [Implemented]

PLAT-005: Raspberry Pi deployments shall support appropriate native camera interfaces, including Raspberry Pi camera modules where compatible, in addition to supported USB cameras.  [Implemented]

PLAT-006: Core scientific processing, calibration, configuration, logging, synthetic spectrum generation, and data analysis functionality shall remain functionally equivalent across Linux, Windows, and Raspberry Pi platforms, subject only to hardware-specific capabilities and performance constraints.  [Planned]

PLAT-007: Platform-specific dependencies shall be isolated and documented.  [In Progress]

# 4. Configuration Management

CFG-001: Configuration shall be externalized from source code.  [Implemented]

CFG-002: INI shall be the preferred configuration-file format.  [Implemented]

CFG-003: Multiple configuration profiles shall be supported.  [Implemented]

CFG-004: Every configuration key/value pair shall be documented with purpose and valid options.  [Implemented]

CFG-005: Invalid configuration values shall produce meaningful diagnostics and graceful termination.  [Implemented]

CFG-006: Command-line options may temporarily override configuration values without modifying the configuration file.  [Implemented]

# 5. Command-Line Interface

CLI-001: The application shall provide a command-line interface.  [Implemented]

CLI-002: Users shall be able to select configuration files and operating profiles.  [Implemented]

CLI-003: A verbose option shall provide additional diagnostic output.  [Implemented]

CLI-004: Diagnostic modes shall support camera testing, hardware verification, calibration checks, configuration validation, synthetic-data generation, and analysis.  [Implemented]

# 6. Logging and Error Handling

LOG-001: The system shall provide structured application logging.  [Implemented]

LOG-002: Logs shall be stored in a dedicated logs directory and support rotation.  [Implemented]

ERR-001: Predictable configuration, hardware, camera, file-system, and initialization failures shall be handled gracefully.  [Implemented]

ERR-002: Fatal errors shall be logged before controlled termination whenever logging is available.  [Implemented]

# 7. Synthetic Spectrum and Test Data Generation

SYN-001: The system shall generate synthetic spectral test data for analysis and regression testing.  [Implemented]

SYN-002: The user shall be able to specify multiple wavelengths.  [Implemented]

SYN-003: The user shall be able to specify line amplitudes or relative intensities.  [Implemented]

SYN-004: The user shall be able to specify FWHM globally and, where supported, per spectral line.  [Implemented]

SYN-005: The user shall be able to specify signal-to-noise ratio.  [Implemented]

SYN-006: The generator shall support wavelength range, sampling interval, baseline level, baseline slope, and deterministic random seeds.  [Implemented]

SYN-007: Gaussian profiles shall be supported initially; Lorentzian and Voigt profiles should be supported as future extensions.  [In Progress]

SYN-008: Synthetic spectra shall pass through the same analysis pipeline used for acquired spectra.  [Implemented]

SYN-009: Generated datasets shall retain complete generation metadata for reproducibility.  [Implemented]

# 8. Hardware and Camera Support

HW-001: Hardware interfaces shall be platform independent where practical.  [Implemented]

HW-002: USB cameras shall be supported where compatible.  [Implemented]

HW-003: Raspberry Pi native camera interfaces shall be supported through dedicated adapters where compatible.  [Implemented]

HW-004: Additional camera models shall be addable without redesigning scientific-processing modules.  [Implemented]

# 9. Data and Scientific Reproducibility

DATA-001: Raw images, processed spectra, calibration data, and measurement metadata shall be savable.  [In Progress]

DATA-002: Stored datasets shall preserve sufficient configuration and calibration metadata to reproduce analysis.  [Implemented]

# 10. Testing and Acceptance

TEST-001: Core scientific behavior shall be tested on Linux, Windows, and Raspberry Pi.  [In Progress]

TEST-002: Synthetic spectra shall be used to verify wavelength recovery, FWHM recovery, amplitude recovery, blended-line behavior, and SNR behavior.  [In Progress]

ACC-001: The application shall run reliably on supported Linux, Windows, and Raspberry Pi systems.  [In Progress]

ACC-002: Core scientific functionality shall remain equivalent across first-class platforms, subject to documented hardware and performance constraints.  [Planned]